scopeblock — Project Plan

## Function List

| # | Function | Key SCPI / basis |
| --- | --- | --- |
| 1 | Initial script (bench_connect.py) | open_resource, *IDN?, *CLS |
| 2 | connect() / identify() | *IDN? |
| 3 | configure(config) | CH / HORizontal / TRIGger / DATa |
| 4 | acquire(timeout) | ACQuire:STATE / *OPC? |
| 5 | read_waveform(source) | WFMOutpre? / CURVe? |
| 6 | measure(wf, kind) | feature extraction |
| 7 | compare(wf, expected, tol) | Result / Tolerance |
| 8 | export(wf, path, fmt) | CSV / npz / MF4 |
| 9 | check_errors() | *ESR? / ALLEV? |

# Function 1: Initial Script — bench_connect.py (Expected By: 7/6/2026)

python scripts/bench_connect.py [--resource RES] [--timeout MS] -> exit code

The first script to be written, scripts/bench_connect.py: a small, standalone pyvisa program that opens a VISA session to the instrument, confirms its identity, and pushes a few commands. It becomes the blueprint for the VISA transport in Function 2.

## ANALYZE — what it does

- Purpose: prove a live, two-way SCPI session to the MSO44B over VISA: open the resource, query *IDN?, quiet the protocol, read the error queue, and close cleanly.

- In between: create a ResourceManager, open the resource, set timeout and read/write terminations, write/query a handful of commands.

- Error modes / edge cases: SCOPE_RESOURCE unset (list visible resources and exit), instrument not found, VISA backend missing, query timeout (no reply).

Parameters

| Parameter | Type / default | Meaning |
| --- | --- | --- |
| --resource | str (default: $SCOPE_RESOURCE) | VISA resource string, e.g. TCPIP0::192.168.0.10::INSTR or USB0::0x0699::...::INSTR. Never hardcoded; the CLI flag overrides the environment variable. |
| --timeout | int ms (default: 5000) | Hard I/O cap so a no-reply raises instead of hanging. |

Outputs

| Output | Where | Meaning |
| --- | --- | --- |
| Identity | stdout | The *IDN? string |
| Error queue | stdout | The ALLEV? contents after hygiene (standard commands) empty on a healthy session. |
| Exit code | process | 0 = session opened, identified, closed clean; 2 = SCOPE_RESOURCE unset (resources listed); non-zero on open/timeout failure. |

## IMPLEMENT — how to build it

- Use pyvisa.ResourceManager() and open with a context manager so the session always closes: with rm.open_resource(res) as inst:.

- Read the resource from the environment; if unset, print rm.list_resources() and exit with guidance — no hardcoded IP or serial.

- Set inst.timeout (a hard cap so a no-reply can never hang), and read_termination / write_termination to '\n'.

- Send session hygiene once: HEADer OFF, VERBose OFF, *CLS — so query replies are bare values and the error queue starts clean.

- Best practice: keep it under 60 lines for simplicity; use argparse to override resource/timeout; do not enable any output or change instrument state beyond the simple commands.

## VERIFY & VALIDATE

- Verify W/out Hardware: run against the pyvisa-sim backend with a small YAML that defines a *IDN? reply, so the script's control flow (open, query, hygiene, close) is exercised with no hardware. Lint (ruff) and type-check (mypy) clean.

- Validate with Hardware: run on the bench with the real SCOPE_RESOURCE; confirm *IDN? returns a Tektronix MSO44B identity, that ALLEV(Error Queue) reports no errors afterward, and that the session closes without a lingering lock.

- Result Expectation: with the bench connected the script prints the correct identity and an empty error queue and exits 0; with SCOPE_RESOURCE unset it lists resources and exits non-zero without traceback.

# Functions 2–9: MSO44B Oscilloscope Library

The oscilloscope library funnels everything through one contract, the Waveform (numpy time/voltage arrays in seconds/volts plus a metadata dict). All SCPI/VISA I/O sits behind a thin Transport interface (write / query / query_binary / close), so every function below is testable offline with a fake transport. The functions follow in build order.

## Function 2:  connect() / identify()  (Expected By: 7/8/2026)

VisaTransport(resource, timeout) ; MSO44B(transport).session_setup() -> str (the *IDN?)

### ANALYZE

Establishes a session and confirms instrument identity. In the library this is two pieces: a VisaTransport that opens the VISA resource and implements the Transport interface, and MSO44B.session_setup() that puts the instrument in a known, parse-friendly state and returns its identity.

| Parameter | Type | Meaning |
| --- | --- | --- |
| resource | str | VISA resource string (from SCOPE_RESOURCE), for the transport. |
| timeout | float (ms) | Hard I/O cap so a no-reply raises instead of hanging ; hard maximum time for the read response |
| transport | Transport | The I/O seam MSO44B drives (VISA on the bench, fake in tests). |

- Returns: the *IDN? string.

- Side effects: writes HEADer OFF, VERBose OFF, *CLS; marks the session ready.

  - HEADer  OFF means when you run float(reply) it would return a clean value

  - VERBose OFF means you are given a short and consistent response

  - Clear status allows you to make sure that any errors that are in the error queue are from this session

- Errors: TransportError on open failure or timeout.

### IMPLEMENT

- VisaTransport wraps pyvisa: write/query (strip whitespace) and query_binary (read the raw IEEE block via read_raw); translate pyvisa timeouts into TransportError.

- session_setup() queries *IDN?, sends the three hygiene commands, sets a _session_ready flag

(so it runs once).

- Best practice: import both pyvisa and build SCPI

### VERIFY & VALIDATE

- Verify: with a FakeTransport returning a fake *IDN?, assert session_setup returns the identity and issues HEADer OFF / VERBose OFF / *CLS in order; assert it is idempotent (a second call does not re-run hygiene).

- Validate: on the bench, assert the identity matches /TEKTRONIX,MSO44/ and the error queue is empty.

- Acceptance: a real session opens, identifies, and reads back clean; the fake-transport unit test passes with no hardware.